# AUTOMOTIVE INDUSTRY STANDARD

Medical Equipment for

Life Support Two Wheeled Road Ambulances

Date of hosting on website: 16th December, 2025

Last date for comments: 14th January, 2026

INTRODUCTION

The Government of India felt the need for a permanent agency to expedite the publication of standards and development of test facilities in parallel when the work on the preparation of the standards is going on, as the development of improved safety critical parts can be undertaken only after the publication of the standard and commissioning of test facilities. To this end, the erstwhile Ministry of Surface Transport (MOST) has constituted a permanent Automotive Industry Standards Committee (AISC) vide order No. RT-11028/11/97-MVL dated September 15, 1997. The standards prepared by AISC will be approved by the permanent CMVR Technical Standing Committee (CMVR- TSC). After approval, the Automotive Research Association of India, (ARAI), Pune, being the Secretariat of the AIS Committee, will publish this standard. For better dissemination of this information ARAI may publish this document on their Web site. During the deliberations, it was decided that type approval of vehicles shall be limited to automotive aspect of construction and functional requirements of life support two wheeled road ambulance and medical equipment’s related details shall not be part of Vehicle Type Approval. This Part 2 of the AIS-209 describes the specific requirements for Medical Equipment that shall be fitted to qualify it as “ambulance”.

TABLE OF CONTENTS

| Clause No | Description | Page No |
| --- | --- | --- |
| 1 | Scope | 04 |
| 2 | Reference Standards | 04 |
| 3 | Definitions | 04 |
| 4 | General Requirements | 04 |
| 5 | Provision for Patients Safety and Medical Equipment’s | 05 |
| Annexure 1 | Information to be submitted at the time of approval | 14 |
| Annexure 2 | Composition of AISC Panel | 16 |
| Annexure 3 | Medical Equipment(s) for Ambulance Accessory Unit | 17 |

Medical Equipment (1) for Life Support Two Wheeled Road Ambulances

Note: Clause from 4.1 to 4.7 are taken from AIS-125 (Part 2)

|  | ### SCOPE |
| --- | --- |
|  | This standard specifies the guidelines regarding technical requirements for medical equipment which shall be installed in life support two-wheeled road ambulances for qualifying it as ‘ambulance’. The concerned RTO to ensure that fitment of medical devices contained in this part in the ambulance accessory unit without compromising safety in consultation with committee of medical experts of the State. |
|  | ### REFERENCE STANDARDS |
|  | AIS-209 (Part 1) Construction and Functional Requirement For Life Support Two Wheeled Road Ambulance |
|  | ### DEFINITIONS |
|  | For the purpose of this standard, definitions given in AIS0-209 (Part 1) shall be referred. The following definition shall apply: |
|  | General Requirements |
| 4.1 | The Equipment should be designed for use in mobile situations and in field applications. If a medical Equipment is designated as "portable", which is meant for use inside an ambulance (except patient handling equipment according to Table 1) it should be in accordance with ISO/IEC 60601-1 and should |
|  | - Be possible to be carried by one person - Have its own built in power supply (where relevant) - Be capable of use outside the vehicle - Be placed preferably along the street side wall of the patient compartment or along the ceiling ensuring the minimum possible distance to be connected to the patient without hindering the movement of personnel around the main stretcher |
| ### 4.2 | Temperature |
|  | - Medical equipment should be complaint with respective product standards as applicable (ISO/ BIS/ US FDA/ CE). AMC/ CMC and Calibration of the equipment shall be verified periodically by ambulance owner. |

| ### 4.2.1 | Unless otherwise marked on the Equipment, the Equipment should function as described in 4.2.2 and 4.2.3 when brought back to room temperature (20°C) after storage in temperatures ranging from -30°C to 70°C. |
| --- | --- |
| ### 4.2.2 | Unless otherwise marked on the Equipment, the Equipment should function throughout the temperature range from 0°C to 40°C. |
| ### 4.2.3 | Unless otherwise marked on the Equipment, the Equipment should function for at least 20 min when placed in an environment at -5°C after storage at room temperature (20°C). |
| ### 4.3 | Humidity and ingress of liquids Equipment’s should comply with ISO/IEC 60601-1 and with particular Equipment standards of the series ISO/IEC 60601-2 where applicable. |
| ### 4.4 | Mechanical strength |
| ### 4.4.1 | General |
|  | Where there are not more stringent requirements for mechanical strength in particular Equipment’s’ standards, then the following mechanical strength requirements should apply to medical Equipment’s for use in road ambulances. |
| ### 4.4.2 | Vibration and bump |
|  | After vibration tests and bump test in accordance with 4.11 the maintain system and Equipment should function within the tolerances specified by the manufacturer. |
| ### 4.4.3 | Free fall |
|  | If the medical Equipment is fixed, as defined in ISO/IEC 60601-1 it is exempted from the free fall test. Equipment’s which are taken out of holders and/or carried by hand should be submitted to the free fall test according 4.12 and should then function within the tolerances specified by the manufacturer. Note: A medical Equipment may consist of fixed and loose components, the free fall test applies to the loose components only. |
| ### 4.5 | Fixation of Equipment’s |
| 4.5.1 | The fixation system(s), maintain system(s) or storage system(s) should hold the Equipment to withstand accelerations or decelerations of 10 g longitudinal (forward, backward), 10 g transverse (left, right) and 10 g vertical. Terminal units and electrical socket outlets should not be used as part of the fixation system. If rails systems are used, they should comply with ISO 19054. Note: Rail systems consist of e.g. rail supports, rails, rail clamps, equipment mount holders, equipment mounts, equipment pin holders and equipment pins. |
| ### 4.6 | Electrical safety |
|  | All Equipment’s should be selected and mounted so that no harmful influence to the electrical supplies results. |
| 4.7 | User interface |
|  | Buttons, switches, indicators and controls should be easily accessible and visible. SI units (except for blood pressure and airway pressure) and standardized graphical symbols where applicable should be used. |
| ## 5.0 | ## PROVISION MEDICAL EQUIPMENTS |
| ## 5.1 | ## Two wheeled Road Ambulance shall have medical equipment’s as given in Table 1 below. |

Table 1

Medical Equipment(s) for Ambulance Accessory Unit (1)

| S.No. | Equipment | Requirement(s) |
| --- | --- | --- |
| 1. | Main Stretcher/ Under Carriage (2) | 01 |
| 2. | Cervical Upper Spinal Immobilization Equipment (2) | 01 |
| 3. | Multipurpose malleable splint with compression bandage set | 01 |
| 4. | Portable oxygen (2) | 01 |
| 5. | Extended Support oxygen (Type B Cylinder) (3) | 01 |
| 6. | Valve for Cylinders 1 | 01 |
| 7. | Resuscitator or with oxygen inlet and masks and airways for all ages and oxygen reservoir | 01 |
| 8. | Mouth to mask ventilator with oxygen inlet | 01 |
| 9. | Electrical or pedal Portable Aspirator & its manual | 01 |
| 10. | Stationary Oxygen (Min 10 LPM oxygen concentrator) (4) | 01 |
| 11. | Pain Relief | X |
| 12. | Infusion Solution 1 Litre | 02 |
| 13. | Material for Treatment of Wounds | 01 |
| 14. | Tourniquet and bleed controlling Equipment’s kit (clotting agents & bandages etc). | 01 |
| 15. | Material for Treatment of Burns and corrosives | 01 |
| 16. | Vomiting Bag | 02 |
| 17. | Non-glass urine bottle | 01 |
| 18. | Sharps container | 01 |
| 19. | Sterile Surgical Gloves, Pairs | 05 |
| 20. | Non-Sterile Gloves for Single Use | 100 |
| 21. | Emergency Delivery Kit (3) | 01 |
| 22. | Waste Bag (5) | 01 |
| 23. | Clinical Waster Bag (color coded yellow, blue and black) | 15 (05 each) |
| 24. | Non-woven stretcher sheet | 01 |
| 25. | PPE Kit (as per standard government norms) for daily use and/or infection control as per site of deployment. | 02 |
| 26. | Cleaning and disinfection material (2) | 01 |
| 27. | Rescue tools | X |
| 28. | Seat belt cutter | 01 |
| 29. | Warning Triangle Lights | 02 |
| 30. | Fire Extinguisher, ABC Type (min 2 Kg capacity) (2) | 01 |
| 31. | Mobile phone dedicated with Ambulance in sync with fitted GPS Equipment | 01 |
| 32. | Internal communication between patient compartment and outside (2) | 01 |

# All the items shall be kept in inventory for continuous replacement for daily/ weekly usage with replacement Equipment’s available ensuring minimum response time during emergencies.

# Specifications and positioning is mentioned in AIS-209

# b Optional as per space availability;

# Essential supply for which fitment is design to use it during emergency while using ambulance accessory unit as standalone unit;

# ‘X’ requirements refers to site of deployment or user’s specific requirement.

# Waste bag to be a Biohazard bag. Waste generated to be disposed as per applicable Biomedical waste handling norms

ANNEXURE 1

(See Introduction)

COMPOSITION OF AISC PANEL ON

CONSTRUCTION AND FUNCTIONAL REQUIREMENT FOR LIFE SAVING-TWO WHEELED ROAD AMBULANCE

| Convener |  |
| --- | --- |
| Mr. Amit Gupta | All India Institute of Medical Sciences |
| Members |  |
| Mr. A. A. Badusha | ARAI |
| Mr. Vikram Tandon | ARAI |
| Mr. Kamalesh B. Patil | ARAI |
| Mr. S. N. Dhole | CIRT |
| Mr. Deepak Joshi | ICAT |
| Mr. Gurkaran Singh | ICAT |
| Mr. Deepak Joshi | ICAT |
| Mr. Nagarajan S | GARC |
| Mr. Raman Chawla | DRDO |
| Mr. K.Kapila | IRF |
| Mr. Balraj Bhanot | IAC & IRF |
| Mr. Parmeshwar Mane | Bajaj Auto Ltd. |
| Mr. Piyush Chowdhry | Hero Moto. Corp. Ltd. (SIAM) |
| Mr. Ishan Gupta | Hero Moto. Corp. Ltd. (SIAM) |
| Mr. Shiva Arun Teja | Hero Moto. Corp. Ltd. (SIAM) |
| Mr. Apoorva Tripathi | Hero Moto. Corp. Ltd. (SIAM) |
| Ms. Garima | Hero Moto. Corp. Ltd. (SIAM) |
| Mr. S. Gururajan | TVS Motor Company Ltd. (SIAM) |
| Mr. Sanchit Khare | OLA Electric |

# ANNEXURE 2

# (See Introduction)

# COMMITTEE COMPOSITION *

# Automotive Industry Standards Committee

| Chairperson |  |
| --- | --- |
| Dr. Reji Mathai | Director, The Automotive Research Association of India |
| Members | Representing |
| Representative from | Ministry of Road Transport and Highways |
| Representative from | Ministry of Heavy Industries |
| Representative from | Office of the Development Commissioner, MSME, Ministry of Micro, Small and Medium Enterprises |
| Shri Shrikant R. Marathe | Former Chairman, AISC |
| Head-TED | Bureau of Indian Standards |
| Director | Central Institute of Road Transport |
| Director | Global Automotive Research Centre |
| Director | International Centre for Automotive Technology |
| Director | Indian Institute of Petroleum |
| Director | National Automotive Test Tracks |
| Director | Vehicles Research and Development Establishment |
| Director | Indian Rubber Manufacturers Research Association |
| Representatives from | Society of Indian Automobile Manufacturers |
| Representatives from | Tractor and Mechanization Association |
| Representatives from | Automotive Components Manufacturers Association of India |
| Representative from | Indian Construction Equipment Manufactures' Association |
| Member Secretary |  |
| Shri Vikram Tandon | The Automotive Research Association of India |